R22 and R27 at 5k causes output to max out around .68 V even with potentiometer calibrating

The last op amp seems to output at a ratio of 7.196 at 500 mv dc input
its ratio is 7.197 at 100mv dc input
its ratio is 7.19 at 10 mv dc input

From input to first opamp output comparison
500mv to 2656 mV
100mv to 531.37 mV
10mv 53.137 mV
Meaning ratio is 5.3137

with -10mV input
3rd opamp output 53.136mV
Long line stays at 53.11
Long line feeds into + input of 5th opamp, also has .1uF capacitor
5th Opamp output is 72.78 mV

Theoretical Calculations:

1st Op Amp: 1+ 110/25.5 = 5.3137 ratio
2nd Op Amp: Full Wave Rectifier
3rd Op Amp: Unity Gain
4th Op Amp: Low Pass Filter?
5th Op Amp: Low Pass Filter? And 1+3.65/10 = 1.365 gain

Meaning DC input voltage to 5th Op Amp Output is: 5.3137*1.365 = 7.2532 gain

Sine Wave LTSPICE Test:

Note: 1st Op Amp may be 1 + 110/(22.5k||187k+13k[reactance of .2uF capacitor]) = 1 + 110k/22.62k = 5.86 Gain
Input of 100mv Amplitude sine wave
1st Op Amp: Output is 587 mV Amplitude
2nd Op Amp: 587 mV (rectified)
3rd Op Amp: 587 mV
Long Line 336.5mV
400mV

1v at input is .748v at 2x
2v at input is 1.38v at 2x
3V at input is 1.38V at 2x
.5v at input is .374 at 2x
Results suggest it is linear when changing the multiplier value of 1st opamp. (R3 from 110k to 25.5k). Should double the input going across the 1st op amp

Working on Solution:

2.82V amplitude (2V rms) produces 16.56V
if I can turn 2.82V input into 5v then solution should work (1.77 gain)
1.77 = 1 + R3/25.5k
R3 = .77*25.5k = 19.635k
To do that I use 16.56V

LTSPICE Results
Right now best results is R3 as 17k and doing this voltage divider for last op amp

Next steps:
I notice with dc values when I run 1V I expect 1V at the 1x output but I get 1.058V. That is a 58mV difference. Compared to the original of 1.987V which is a 11 mV difference. I hypothesize the difference is because I tuned the resistors in the circuit for sine wave outputs rather than dc. Do to impedance issues there are probably some differences. One way I think to change it is to mess with the first op amp r5 and c1 values.

First test. Default values of 187k and .1uF gets max value of 5.2456. Calculating the impedance, I get the value of 188.8k, 188.8k in parallel with 25.5k is 22.47